## Slide 1: IN SITU ADCP
Notes: 1

## Slide 2
ADCP IN SITU Features
ADCP IN SITU is used to process data produced by profiling current meters:
Instantaneous pressure measurements (on certain instruments);
Instantaneous temperature measurements;
Heading, roll and pitch measurements;
Current measurements integrated over a period (integration period);
Measurement of the back-scattered echo or of the amplitude of the received signal according to the instrument in question.
ADCP = Acoustic Doppler Current Profiler.

## Slide 3
Reads the raw data coming from Sontek, Nortek and RDI equipment.
  Sontek
    *.adp format
  Nortek NDP
    *.adp format
  Nortek Aquapro
    *.prf format
  Nortek Aquadopp (single point current meter)
    *.aqd format
  Nortek Awac
    *.wpr format
  Nortek continental
    *.cpr format
  Teledyne-RDI
    *.000 format
Input Data

## Slide 4
Setting the Import Parameters

## Slide 5
Data Display and Export

## Slide 6
Accessing ADCP In Situ

## Slide 7
Description of ADCP IN SITU
All the functions are accessible from a control window.
When you first save data from the In Situ Acoustic Doppler viewer module, an ADOP directory is created under the C:\HYPACK 20XX\Projects\[name of the project].
The *.hcf (project colors) file manages the various color palettes required for the final display will be stored within it .
Predefined *.hcf files under C:\HYPACK 20XX\ADCP
7

## Slide 8
ADCP In Situ Shell
 | Setup | Raw file import parameters.
 | File Open | Load ADCP, Log or equipment identification file.
 | File Name | File currently being processed. Use arrows to scroll through files loaded from a LOG.
 | Information | Information about the raw file loaded.
 | Data Tables | Access data spreadsheet.
 | Profile | Access data graphic display.
 | Save | Raw data in HYPACK *.edd file.
 | Export | Data to ASCII, Etude, Tide or ODV format Equipment identification file
 | Help | Access Help files.

## Slide 9
Import Parameters
After selecting a raw data file, click on the Setup icon in order to set the import parameters.

## Slide 10
Station Settings
Id: equipment serial number.
Latitude and Longitude: mooring position (for exporting in Etude format, the format must be DDDMM.MM)
Hydrographic Depth: depth with respect to the hydrographic datum.
Instrument Depth: submersion depth of the current meter i.e. Hydrographic Depth – sensor offset
Instrument Orientation: mounting of the sensor facing upwards or downwards.
Tide File: selection of the tide file in HYPACK (*.tdx) or Masg (*.tdf) format.

## Slide 11
Output Tab
Water layer that you wish to export in the Etude format.
There are two ways in which the reference of the exported water layers can be defined:
  Select Measure at Hydrographic Depth: the reference will be the hydrographic datum (the reference for exporting will be the surface).
  Do not select Measure at Hydrographic Depth: the reference will be the water level, i.e. the surface detected according to the method chosen. A tide file must be associated.

## Slide 12
Referencing Hydrographic Datum
Select `Measure at Hydrographic Depth’.
The current for exporting is therefore calculated between 10 and 30 m.

## Slide 13
Referencing the Water Level
Do not select `Measure at Hydrographic Depth’.
The current calculation is carried out in the layer 30-40 m from the surface.

## Slide 14
The Measure at Hydrographic Depth option is not selected…
Referencing the Water Level

## Slide 15
Time Tab
Time Zone Shift: Correction of the time zone. Indicate the time in minutes.
Time drift: Correction of the clock drift, indicate the time in seconds.
  The drift correction is applied linearly across the entire period. It represents the total drift.
  The drift is positive if the equipment’s clock time increments more quickly than the actual time.
Time shift: Shifts the time to the middle of the integration period. Indicate the time in seconds.
The module makes this correction automatically for Sontek, Nortek and RDI data..

## Slide 16
Complements about time
The module makes this correction automatically for Sontek, Nortek and RDI data..
Notes: 16

## Slide 17
Pressure Tab
The processing is carried out as follows:
Calibration correction.
Calculation of the actual depth of the instrument by integrating the pressure data. The depth is calculated in meters using the UNESCO formula.
Manual validation with graphic display.

## Slide 18
Temperature Tab
Calibration: The sensor calibration coefficient is used to correct a series of values by applying a polynomial function.
The default values (a0 = 0 and a1 = 1) result in no data adjustments calculated.
Corrected Value = A0+(M* A1) where M = the measured value.

## Slide 19
Salinity Tab
Calibration:
Currently unused.
The equipment has no salinity measurement sensor.
The a0 field can be used when the data files have no pre-defined salinity value to complete this value.
User-defined in the Nortek configuration software and used to calculate speed of sound. The salinity calibration constants can be used to approximately correct an error in the programmed salinity value.

## Slide 20
Speed of Sound Tab
Calibration:
  Currently unused.
  The equipment has no speed of sound measurement sensor.
Disable Del Grosso Correction: This option allows you to use a speed of sound fixed at 1500 m/s for all values in place of a speed of sound correction.

## Slide 21
Attitude Tab
Magnetic declination: Correction of the magnetic declination (i.e. angle between the magnetic north and true north), this value is positive to the East and negative to the West, value in degrees.
Magnetic deviation: Correction of the compass deviation (interference if the current meter is moored in a weighted metal cage), the convention and units are the same as for declination.
Heading, Pitch, Roll: Sensor calibration coefficient used to correct series of values by applying a polynomial function.
Corrected Value (Polynomial) = A+(M+ B)+(M2+ C)+(M3+ D)+(M4+ E)+(M5+ F)

## Slide 22
Saving the Equipment Identification File
You can save the equipment configuration file specific to your equipment or a project before starting to process the data.
  Click on the hot key.
  Select saving with the extension*.ini
Reload the configuration later using the File Open icon, if necessary.

## Slide 23
Data Displays
The display of data including the parameters for each sensor is shown by clicking on the Data Tables icon.

## Slide 24
Data Displays

## Slide 25
Data Displays
Data pertaining to each condition monitored.
Each spreadsheet shows:
  raw data,
  data corrected according to the Calibration settings.
Right-click to access graph options.
  Select the value you want in graph view.

## Slide 26
Validating Data in the Spreadsheets
To invalidate values:
Remove the checks in front of each record.
Make a multiple selection (SHIFT key) then pressing the DEL key on the keyboard.
In order to cancel the invalidation of a value:
Click on the ticks in front of the invalidated recordings.
Make a multiple selection (CTRL key) of the invalidated recordings then press the DEL key on the keyboard.
To validate all the lines (whether or not the values are ticked):
Make a multiple selection.
Press the ESC key.
Note: If the data graph is open, you must update it for your spreadsheet edits. Right-click and select Refresh Graph.

## Slide 27
Validating Data in the Graphs
Right-click spreadsheet to access graph options.
To graphically select data: Press the CTRL key and click-and-drag a box with the cursor. The data selected appears marked in yellow:
  To cancel the selection, press the ESC key.
To invalidate selected data press the DEL button on the keyboard.
To cancel the invalidation of values assigned graphically, select the values in the spreadsheet and press DEL again.
View Adjustment | Cursor Action
Zoom window | Click and drag a box from its top left to its bottom right corner.
Zoom Extents | Click and drag a box from bottom right to upper left.
Zoom In/Out | Use your mouse wheel.

## Slide 28
Right click on the spreadsheet, you can access a graphic representation of the values:
Graphs

## Slide 29
Filtering Data
This dialog box is used to apply filters to all the following data:
  Pressure
  Temperature
  Salinity
  Heading
  Roll
  Pitch
  Scale, filter on the recording counter.
The filter for each type of value is set by changing tab.

## Slide 30
Processing and Display of Current Data

## Slide 31
Profile Display
Operating data are displayed by clicking on the Profile icon.

## Slide 32
Profile Display Settings
Click the Setup Icon.
Options:
  Color Palettes
  Graphic Graduation
  Pressure Sensor Display
  Tide Value

## Slide 33
Icon Functions

## Slide 34
Profile Display Options
 | Print graphic displayed.
 | Save display to bitmap image.
 | Zoom window
 | Zoom Extents
 | Access Filters dialog
 | Undo/Redo the last filter operation
 | Delete data above the interface detection line. Data is also deleted from the export and graphic displays.
 | Data Viewer displays the data at the cursor position.

## Slide 35
This double entry function is used to select the items displayed.
A principle value can be displayed with various components .
Profile Display
Velocity :
  East
  North
  Vertical
  Error
  Magnitude
  Direction
Correlation
  Beam 1, 2, 3 and 4
  Error
Amplitude
  Beam 1, 2, 3 and 4
  Error
Pct Good
  PG1, 2 , 3 and 4

## Slide 36
Profile Color Settings
Independent color settings for each component.

## Slide 37
Data Viewer
Data Viewer displays the data indicated by the cursor.

## Slide 38
Profile and Time Series View
Display of graphic data by profile or time series

## Slide 39
Profile and Time Series View
Display of data by direction. By profile or time series.

## Slide 40
Functions

## Slide 41: Particularity of single point current meter
The single point current meter Nortek Aquadopp makes measure in one cells, you’ve got one profile of one cells.
The processing of auxiliary data will be the same that profiling current meter.
The data visualization will be little bit different, the profile window displaying one cells, the cell number 1.

## Slide 42
Validation / Invalidation of Graphic Data
Two methods can be used to invalidate values:
  Use of the filters.
  Use of the cursor.

## Slide 43
Filtering Data
Click the Filter icon.
The following values can be filtered:
  -Velocity (East, North, Up, Error).
  -Amplitude (Beam 1, 2, 3 and 4).
  -Correlation (Beam 1, 2, 3 and 4).
  -Percentage, Pct good (PG1, 2, 3 and 4).
  -Magnitude.
Activate the filter by checking the Enable box.

## Slide 44
Validating Data in the Profile Display
There are three ways to invalidate data in the Profile window:
  By cell.
  By profile.
  By level.
This parameter is defined in the profile Setup window.

## Slide 45
Validating Data in the Profile Display
Graphic data can be invalidated using the cursor.
To activate the cursor tool, check Enable Cursor in the Profile window:
To Invalidate data, point to the data in the profile to be invalidated and press Delete on the keyboard. The data will be replaced by white pixels.
To re-validate data, point to the previously invalidated data using the cursor and press ESC on the keyboard.
Profile Cell Level

## Slide 46
ADCP Player
View Adjustment | Keyboard
Zoom window | Click and drag a box from its top left to its bottom right corner.
Zoom Extents | Click and drag a box from bottom right to upper left.
Zoom In/Out | Use your mouse wheel +/- keys on keyboard
Increase/Decrease Magnitude Scale | Page up/down
Change Horizontal Angle | Up/Down Arrows
Change Rotation Axis | Left/Right Arrows
Return to first recording | Esc key
Control Panel Display of Reading Time
Animated display of the magnitude values as a function of time.

## Slide 47
Saving Edited ADCP Data
Click on the Save icon to save the file to *.edd format in the project’s adop\edit directory.
The settings parameters are saved in *.edd format
All the parameters and processing are saved.

## Slide 48
Exporting Your Data
Export your data to various formats by clicking the Export Out icon.
  Tide file
  Etude file
  ASCII format
  ODV format
  Equipment configuration file
Select the export format from the File of Type drop-down in the Windows Save dialog.
Name the file to be generated and click [Save].

## Slide 49
Exporting Tide Data
The interface can be exported in the HYPACK *.TDX or MASG *.TDF format.
  Click the Export Tide Icon and name your file.

## Slide 50
Exporting Tide Data
The interface can be exported in the HYPACK *.TDX or MASG *.TDF format.
  Click the Export Tide Icon and name your file.
The data exported depends on your choice of interface detection method.

## Slide 51
Exporting Tide Data
Display Surface Based on Amplitude
Display Surface Based on Pressure
Notes: 51

## Slide 52
Exporting Tide Data
The tide can be exported in the HYPACK *.TDX or MASG *.TDF format.
The data exported depends on your choice of interface detection method.

## Slide 53
Etude Format
Exporting in Etude format (*.B) uses the data averaging algorithm between two depths as defined in the output tab in the setting window.
The water layer that you wish to export must be a modulo of the size of the cell parameter set on acquiring the equipment.

## Slide 54
ASCII Export
An ASCII export is available. All or the parameters can be set by the operator.
In order to access the ASCII Out Setup window, you must first enter the file name and click [Save] in the File Save dialog.
Select the speed unit.
To choose the values to export use the arrows to move them between the list of available files and the central zone.
  You can also select the cells to be exported.
    Select `Cell No.’ and check the required cell numbers.
The number of decimals for each value is independent.
Customizable fields can be created.
The limits of the fields can be modified.
To customize a file heading, click [Hdg Edit].
To save the export parameters in a configuration file, click [Save Cfg]. This export can be loaded subsequently using the [Load Cfg].

## Slide 55
ODV Export
Export in ODV format from the ASCII export by loading the previously created configuration file odv.ini, available in Hypack 201X/ADCP.
The ODV software configuration file for importing the HYPACK export is also available in HYPACK 201X/ADCP. It is the Config_type_odv.cfg file.

## Slide 56: Thank You !
January 2020
HYPACK on Youtube.com ( Historical Sessions ) HYPACK Live Chat
HYPACK on Youtube.com ( Newer Sessions ) HYPACK Ustream
HYPACK SUPPORT Site HYPACK Website
Links to more information:
Contact Us:
Sales@HYPACK.com
Help@HYPACK.com
(860) 635 - 1500