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it is not possible to switch on the TitroLine alpha and the display stays dark?
The „Memory-Card“ must be built in.
Is the power unit connected?
Chapter 2.1

it is difficult to read the display?
The angle of view of the display can be adjusted. In the menu <ROLL UP> or <ROLL DOWN> the contrast can be adjusted.
Chapter 5.1.1

the size of the unit is not the same as shown in the display?
The size of the unit must be configured with the keys
<ROLL UP> or <ROLL DOWN> or configured in the
<INIT> menu under <SYS>.
Chapter 2.2

Chapter 5.5.6

it is necessary to change the unit?
Open the retention ring, remove the dosing unit, slide it to
the valve support and remove both by pulling them upwards.
Chapter 2.2

the unit is empty and reagents must be filled in?
Directly after switching on the titrator, it is possible to rinse
by pressing the <FILL> key on the keyboard.
Chapter 2.3
Change dosing unit
Chapter 5.5.6.3

the first filling does not work?
Check the screwing and fix it. Then rinse the dosing unit.
Chapter 2.3
Chapter 5.5.6.3

no working method is available after switching on?
The titrator is new. No standard method is copied
into the working menu, no method is created.
see Short Manual
Chapter 5.4.4

it is not possible to change a standard method after having it copied?
The changing of methods is locked (standard configuration). This locking must be changed in the menu <SYS> with <INI> <EDIT STD METHOD>. Select <ON>.
Chapter 5.5.6.7

it is not possible to copy
standard method into the working memory?
There is already a method in the selected memory.
The message „method still exists“ appears. In the case
of 8 existing methods it is necessary to delete a method.
Chapter 5.4.4

Delete method
Chapter 5.5.2.2

air bubbles are coming out of the titration tip
Check the screwings and fix them. Then rinse the dosing unit.
Chapter 2.3
Chapter 5.5.6.3


What is to do if...
User Manual

the titration stops too early?
The criteria for stopping the titration is wrong. In case of criteria <EQ> the sensitivity is too strong, so that the equivalence point is recognized too early.
Chapter 5.4.1.2

in the case of an endpoint titration the titration takes too much time or lasts too long?
The wrong type of jump was chosen.
See table „type of titration“ EP.
Chapter 5.4.1.2

the result of the calculation
is wrong?
Check consistent calculation formula with factors of the
titration. Check under „MODE“ and <<CALC>> the formula and the factors. In the case of weight or sample volume take care that it will be multiplied with the Q factor.
Chapter 5.4.2.1

there is no printout after
the titration?
Is the printer connection o.k.?
Check in the menu „MODE“ <<PRN>> if the cursor is in the right position under <<Centronics>>. Is the printer „online“?
Chapter 5.4.6

the graphic printout shows ASCII signs instead of a
graphic?
The printer does not work in the EPSON or IBM-proprinter mode. Changing of the mode is described in the printer manual.
Check the printer cable.


no balance weights are
available after connecting
a balance? 
Check the correct cable.
Check if the communication parameter of the serial port
of titrator and balance are identical.
Check in “INIT“ menu under <<RS>>.
Check in balance.
In the actual method the option <<balance>> / <<auto>> is not selected.
Chapter 2.4

Chapter 5.5.4

Chapter 5.4.1.1

Chapter 5.7

it is necessary to change the balance weights?
Press <EDIT PARAM.> in the „Start“ menu.
Now the balance data can be edited.
Chapter 5.7

the speed of the Karl-Fischer titration is not o.k.?
Select the correct control of the titration.
The titration control differs between one component system and two component systems, low and high amounts of water.
Chapter 5.4.1.4

the speed or control of the
pH-stat titration is not o.k.?
Select the correct control of the titration. The control depends on the reaction speed and the buffering of the system.
Chapter 5.4.1.5
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