TitroLine alpha
Overview of Method Development
[image: image3.wmf]
TitroLine alpha
Overview of Method Development
[image: image2.wmf]

Functions of the keyboard

PC- / MF-2-keyboard
Keyboard
TZ 2825
Operation
with the TitroLine alpha

<F1>
<START>
Start of the displayed method.

<F2>
<STOP>
Cancellation of the method.

<F3>
<EDIT PARAM.>
Balance data editing.

<F4>
<FILL>
Filling of the cylinder; with <<CALC>> activation and deactivation of transfer parameters.

<F5>
<STORE PARAM.>
After pressing this key the final position of parameterization is reached and the following display appears <<check/store>>. This menu item can be used to decide whether to store or not.

<F6>
<MODE>
Display of the <MODE> survey menu to parameterize a method.

<F7>
<INIT>
Display of the <INIT> survey menu for initial parameterization.

<F8>
<CAL>
Display of the <CAL> menu for the calibration of the measurement channels.

<F9>
+ / -
Change in sign at <<titration end>> mV and input blank value and factors.

<ESC>
<ESC>
<ESC> returns you to the previous level.

ROLL DOWN
ROLL DOWN
Within a particular menu,
the parameters are scrolled through backwards (scroll forward: <ENTER>). Within the menu items: scrolling through the parameters.

ROLL UP
ROLL UP
Within the menu items: scrolling through the parameters.

< ( >
< ( >
< ( >
< ( >
Scrolling within the survey menu to select menu items. Within the menu items:
Moving the cursor to select the parameters. Selection of the standard parameters on the <<STD>> menu.




PC- / MF-2-keyboard
Keyboard
TZ 2825
Operation
with the TitroLine alpha

0...9,
Decimal point
0...9,
Decimal point
Input of method numbers,
sample numbers and parameters.

<ENTER> or
< ( >
<ENTER>
Confirmation of the input parameters.
In the present operating instructions,
both keys are referred to <ENTER>.

<DEL> or
< ( >
<DEL>
Delete an input digit/character
under which the cursor is positioned.

Letters
Letters and ALPHA key
Alphanumeric input is possible.

All other keys
(((
No function.

Adjustment of the Display

In menu ”INIT” push buttons <roll up> / < roll down>!

Titration parameter

To change the pre-set titration parameter, enter the menu ”INIT”.
Under <Sys> switch on the edit of standard parameter. After doing this it is possible to change all parameters.

Titration parameter consist of:

1.
End criteria for the titration

measuring value, ml-consumption, equivalence point, end point
2.
Speed of the titration

drift control, fixed delay
3.
Control of the titration steps

dynamic, linear

Endpoint titration

Titration up to a  defined pH or mV-value. The consumption at this point will be evaluated. Strong jumps will be controlled differently in comparison to weak ones.

Correct adaptation to the slope
(
precise result





Curve type
Example
Risks if a wrong type is selected

1
steep
HCI versus NaOH
overtitration

2
medium
Acetic Acid versus NaOH
( 3:
overtitration
too slow

3
weak
Acetic Acid versus NH4OH
too slow

4
buffered
strong carbonised water
too slow

5
special
Without automatic titration a linear titration

Detection of the equivalence point of a titration

To stop a titration directly after an equivalence point, it is necessary to define the sensitivity of the detection.


High number   =   strong jump   =   insensitive

Attention:
1.
High sensitivity = it is possible that titration will be stopped too early
early
2.
It is recommendable to combine the end criteria with another

criterion (i.e. measuring value or ml-consumption).

Step
Curve shape
Example

1
very weak jump
acetic acid next to hydrochloric acid

2
weak jump
chloride in trace range, Ca/Mg with ISE (ion sensitive electrode)

3
medium jump
phosphoric acid, both jumps

4
normal intense jump
HCI/NaOH

5
very intense jump
Redox-titration with very intense jump -CSB-

Control of the titration

The control of the titration consists of:

Speed of titration

Fixed
After  each dosing step TitroLine alpha  is waiting


for a defined time in seconds.


Enter time in seconds

Drift controlled
TitroLine alpha is waiting until the measuring value


is stable


Enter fast, medium, exact



Duration of the titration:


fast
2-3 min


medium:
3-5 min


exact:
> 5 min


Unstable measuring values ( long duration

Steps of the titration

Linear:
constant dosing steps. Enter in ml

Dynamic:
adaptation to the slope:


steep slope
(
small steps


flat slope
(
big steps

Attention:
most curves will be titrated with step 1

Dynamic steps

Dynamicstep
Application
Example

1
Usual pH titration with
medium and strong jumps
many mV titrations
HCI with NaOH

chloride with AgNO3

2
Titrations with medium
or several jumps
HCI an acetic acid with
NaOH or chloride and
bromide with AgNO3,
CI in small quantities,
iodometric titration

3
Flat jumps at pH titrations
(at special applications)


4
Explicit statement at some SGH applications



Karl-Fischer Control

The Karl-Fischer determination must be adapted to:

-
consumption
-
speed of water giving off
-
solvent

Reg. step
Application

1
Normal water content / unitary system

2
Low water content / unitary system

3
Normal water content but increased water release;
unitary system and two-component system

4
Normal water content / two-component system

5
Low water content / two-component system

Control pH-stat-titrations


pH Constancy
pH Constancy

Control
stage
 pH < 0,1
 pH < 0,2

1
max. 0,05 mMol H+-
ions/min
max. 0,1 mMol H+-
ions/min

2
max. 0,2  mMol H+-
ions/min
max. 0,5 mMol H+-
ions/min

3
max. 0,5  mMol H+-
ions/min
max. 1,2 mMol H+-
ions/min

4
max. 1     mMol H+-
ions/min
max. 2,5 mMol H+-
ions/min

5
max. 2,5  mMol H+-
ions/min
max. 5,0 mMol H+-
ions/min

Stage 0 is used to define the smallest proton quantities in unbuffered systems.

In case of stage 0 a 10 ml dosing unit has to be used in order to be able to dose the required volumes.


Formulae

ml
=
ml; here calculation is continued with determined ml values
Epn
=
EP1 to EP2; the respectively found end point in terms of ml
Eqn
=
EQ1 to EQ2; the respectively calculated equivalence point


in terms of ml
F
=
Factor 1  (multiplicative)
F2
=
Factor 2  (multiplicative)
C
=
Factor 3  (multiplicative)
Q
=
Quotient; is automatically multiplied by the respective


weighted-in quantity or preset weight in the calculation
B
=
Blank value (additive) 

The formulae are as follows (example for an endpoint titration):

Formula 1:
EPn ( F ( F2 / Q
Formula 2:
(EPn - B) ( F ( F2 / Q
Formula 3:
(EPn+1 - EPn) ( F ( F2 / Q
Formula 4:
(B - EPn) ( F ( F2 / Q
Formula 5:
(Q ( F) / (EPn ( F2)
Formula 6:
(F ( F2) / Epn
Formula 7:
[(B ( C) - (EPn ( F)] ( F2 / Q
Formula 8:
EPn ( F ( F2

Note
:
Formula 3 can only be selected,


if the number of EP/EQs is greater than 1.


Formula 5 can only be selected if EQ = 0.

Weighed-in:
If a weighed-in or sample volume data is entered


the result will always be divided with this value!


Factor Q is an additional multiplication constant.

[image: image1.wmf]
SCHOTT GLAS
Geschäftsbereich Industrieglas
Postfach 24 80
Laborgeschäft
55014 Mainz
Hattenbergstraße 10
Telefon
(06131) 66 49 07
55122 Mainz
Telefax
(06131) 66 40 51

Hersteller:
SCHOTT-GERÄTE GmbH
Im Langgewann 5
65719 Hofheim
Postfach 1130
65701 Hofheim

MEW 000519 E.doc[image: image2.wmf]

[image: image3.wmf]